## Development notes
Developed May-June 2022 by Paul Russo

## Instructions
Purpose:
This workbook permits us to normalize detectors to a standard of finite size.
Using the normalization, you can then push ahead to measure Rg using Guinier
Principle:
There is no such thing as a perfect Rayleigh scatterer. Typically, we might use
a small latex particle, bovine serum albumen, or Brij-97 microemulsion.
The idea is to use something small that scatters a lot, well above solvent level and any
stray light.
The instrument has two kinds of detectors: Wyatt pin photodiodes plus
avalanche photodiode light sensors.
The Wyatt detectors can most easily be read from the front panel display.
More accurate readings can be obtained by collecting data using the Astra software (requires Win98)
The unit of measurement is probably in volts, but it does not matter.
The APD detectors measure in kHz, a totally different thing.
The objective is to figure out what factor puts the measured intensities on the
correct P(q) curve. We calculate the curve from the low-order P(q) expansion, 1-q2Rg2/3.
That will be good enough for any sensible normalization standard, but the equation
could be replaced with a detailed expression for, say, a larger latex sphere.
To use:
1 | Make a DLS run on your normalization standard, save as .ASC file
2 | Normally, this will give you a cumulants Gamma value and Rh value.
3 | Drop those data into the green-shaded cells at the top right; the results are copied where they need to go.
4 | Read the ALV's red panel display for all the laser and angles by gently pressing the selection button; type the results into the appropriate cells (color-coded).
button; type the results into the appropriate cells.
5 | Repeat the previous 2 steps for the solvent (likely a much weaker scatterer).
6 | Enter the Rh value from the DLS run.
7 | Enter your name and date.
8 | Enter the type of normalization standard.
9 | At this point, your normalization constants should be ready.
10 | You can then measure something known, preferably in a solvent with the same refractive index (subtract the solvent).
You can include the Rh value from DLS, but be careful about single-angle DLS (see Gather program for multiangle)
Apply the normalization factors to the intensities.
A Guinier plot or similar should give you the size expected for the known (if it's a sphere, don't forget to convert the DLS results to Rg).
11 | Then you can get Rg/Rh for shape analysis (e.g., should be sqrt(3/5) for solid, uniform spheres)
Testing:
"Circular logic" tests (normalizing with Brij microemulsion, then feeding the same data as an unknown) suggest the normalization factors were determined well enough to feed back into
the measured intensities and come up with smallish errors that might trace to the validity of Rh.
measurement in DLS (i.e., given the viscosity of the microemulsion, the true Rh of Brij blobs might be
smaller than measured.
Thus, Brij is probably perfect for getting the fudge factors right, but agreement between SLS and DLS might be imperfect.

## Normalize
Date of normalization |  | 2022-06-22
Normalizing standard |  | Brij microemulsion
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | y | 0.003 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 4.76 | nm
| p2 |  |  | 13.57 |  | Rg (sphere) | 3.6871 | nm
| p3 |  |  | -14.65 |  | lo/(n*DLS radius) | 108.2012
| LC |  |  | 80 |  | qmaxRg | 0.086256
| L |  |  | 4.93 |  | 1-(qmaxRg)2/3 | 0.99752
Normalization
Comment¯ | Port # | Angle/degrees |  | ---- | Normalizing solution |  | Solvent background |  | Signal | P(q) | Factor
| Zero-ref | 0 |  | Detector type | Wyatt Reading | ALV Reading/kHz | Wyatt Reading | ALV Reading |  | 1 | 1 | Drop solution count rates |  | Drop solvent count rates
| 1 | 22.5 |  | Wyatt diode | 1.055 |  | 0.099 |  | 0.956 | 0.9999 | 1.0459 | here from .ASC file |  | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.71 |  | 0.03 |  | 0.68 | 0.99984 | 1.4704 | MeanCR1 [kHz] : | 740.7855 | MeanCR1 [kHz] : | 3.554
| 3 | 32 |  | Wyatt diode | 0.89 |  | 0.012 |  | 0.878 | 0.99979 | 1.1387 | MeanCR2 [kHz] : | 673.7016 | MeanCR2 [kHz] : | 2.259
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | --- | MeanCR3 [kHz] : | 670.2921 | MeanCR3 [kHz] : | 3.7147
| 5 | 44 |  | Fiber/APD |  | 740.7855 |  | 3.554 | 737.2315 | 0.9996 | 0.0013559 | MeanCR4 [kHz] : | 682.3948 | MeanCR4 [kHz] : | 10.292
| 6 | 50 |  | Fiber/APD |  | 673.7016 |  | 2.259 | 671.4426 | 0.99949 | 0.0014886 | MeanCR5 [kHz] : | 158.7783 | MeanCR5 [kHz] : | 2.6679
| 7 | 57 |  | Wyatt diode | 1.225 |  | 0.034 |  | 1.191 | 0.99933 | 0.83907 | MeanCR6 [kHz] : | 335.1978 | MeanCR6 [kHz] : | 2.7578
| 8 | 64 |  | Fiber/APD |  | 670.2921 |  | 3.7147 | 666.5774 | 0.99916 | 0.0014989 | MeanCR7 [kHz] : | 343.3247 | MeanCR7 [kHz] : | 4.8051
| 9 | 72 |  | Wyatt diode | 1.113 |  | 0.012 |  | 1.101 | 0.99894 | 0.9073 | MeanCR8 [kHz] : | 594.8573 | MeanCR8 [kHz] : | 5.1015
| 10 | 81 |  | Fiber/APD |  | 682.3948 |  | 10.292 | 672.1029 | 0.99865 | 0.0014859
| 11 | 90 |  | Wyatt diode | 0.998 |  | 0.009 |  | 0.989 | 0.99834 | 1.0094
| 12 | 99 |  | Fiber/APD |  | 158.7783 |  | 2.6679 | 156.1104 | 0.99799 | 0.0063928
| 13 | 108 |  | Empty | --- |  | --- |  | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 335.1978 |  | 2.7578 | 332.44 | 0.99719 | 0.0029996
| 15 | 126 |  | Wyatt diode | 1.28 |  | 0.017 |  | 1.263 | 0.99674 | 0.78918
| 16 | 134 |  | Fiber/APD |  | 343.3247 |  | 4.8051 | 338.5196 | 0.99631 | 0.0029431
| 17 | 141 |  | Wyatt diode | 0.705 |  | 0.012 |  | 0.693 | 0.99592 | 1.4371
| 18 | 147 |  | Fiber/APD |  | 594.8573 |  | 5.1015 | 589.7558 | 0.99556 | 0.0016881

## Measure Unknown
Date of measurement |  | June 22 2022
Sample |  | Latex_3A_1:20 dilution
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | y | 0.001 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 92 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 71.2629 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 5.5982
| LC |  |  | 79.2 |  | qmaxRg | 1.6671
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.073551
Normalized
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal | Intensity
| Zero-ref | 0 |  | Detector type | Wyatt Reading | ALV Reading/kHz | Wyatt Reading | ALV Reading |  | 1 | Drop solution count rates |  | Drop solvent count rates
| 1 | 22.5 |  | Wyatt diode | 11.91 |  | 0.124 |  | 11.786 | 12.3272 | here from .ASC file |  | here from .ASC file
| 2 | 28 |  | Wyatt diode | 7.65 |  | 0.053 |  | 7.597 | 11.1703 | MeanCR1 [kHz] : | 5630.6103 | MeanCR1 [kHz] : | 3.554
| 3 | 32 |  | Wyatt diode | 9.17 |  | 0.011 |  | 9.159 | 10.4295 | MeanCR2 [kHz] : | 4565.8354 | MeanCR2 [kHz] : | 2.259
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 4304.6235 | MeanCR3 [kHz] : | 3.7147
| 5 | 44 |  | Fiber/APD |  | 5630.6103 |  | 3.554 | 5627.0564 | 7.6296 | MeanCR4 [kHz] : | 3677.9673 | MeanCR4 [kHz] : | 10.292
| 6 | 50 |  | Fiber/APD |  | 4565.8354 |  | 2.259 | 4563.5765 | 6.7932 | MeanCR5 [kHz] : | 793.5443 | MeanCR5 [kHz] : | 2.6679
| 7 | 57 |  | Wyatt diode | 10.25 |  | 0.02 |  | 10.23 | 8.5837 | MeanCR6 [kHz] : | 1332.102 | MeanCR6 [kHz] : | 2.7578
| 8 | 64 |  | Fiber/APD |  | 4304.6235 |  | 3.7147 | 4300.9088 | 6.4468 | MeanCR7 [kHz] : | 1240.9502 | MeanCR7 [kHz] : | 4.8051
| 9 | 72 |  | Wyatt diode | 8.04 |  | 0.009 |  | 8.031 | 7.2865 | MeanCR8 [kHz] : | 1805.8501 | MeanCR8 [kHz] : | 5.1015
| 10 | 81 |  | Fiber/APD |  | 3677.9673 |  | 10.292 | 3667.6753 | 5.4497
| 11 | 90 |  | Wyatt diode | 5.85 |  | 0.007 |  | 5.843 | 5.8982
| 12 | 99 |  | Fiber/APD |  | 793.5443 |  | 2.6679 | 790.8764 | 5.0559
| 13 | 108 |  | Empty | --- |  | --- |  | --- | ---
| 14 | 117 |  | Fiber/APD |  | 1332.102 |  | 2.7578 | 1329.3442 | 3.9875
| 15 | 126 |  | Wyatt diode | 5.2 |  | 0.013 |  | 5.187 | 4.0935
| 16 | 134 |  | Fiber/APD |  | 1240.9502 |  | 4.8051 | 1236.1451 | 3.6381
| 17 | 141 |  | Wyatt diode | 2.45 |  | 0.006 |  | 2.444 | 3.5123
| 18 | 147 |  | Fiber/APD |  | 1805.8501 |  | 5.1015 | 1800.7486 | 3.0398

## Guinier+NonlinSphere
Normalized
Port | Angle/degrees | q/cm-1 | q2/1010cm-2 | q2/10-4nm-2 | intensity, I | ln (I)
1 | 22.5 | 47599.9678 | 0.22658 | 0.22658 | 12.3272 | 1.0909
2 | 28 | 59026.3747 | 0.34841 | 0.34841 | 11.1703 | 1.0481
3 | 32 | 67252.5891 | 0.45229 | 0.45229 | 10.4295 | 1.0183
5 | 44 | 91400.0326 | 0.8354 | 0.8354 | 7.6296 | 0.8825
6 | 50 | 103114.37 | 1.0633 | 1.0633 | 6.7932 | 0.83207
7 | 57 | 116421.6728 | 1.3554 | 1.3554 | 8.5837 | 0.93367
8 | 64 | 129294.6757 | 1.6717 | 1.6717 | 6.4468 | 0.80934
9 | 72 | 143413.3626 | 2.0567 | 2.0567 | 7.2865 | 0.86252
10 | 81 | 158458.4303 | 2.5109 | 2.5109 | 5.4497 | 0.73637
11 | 90 | 172526.5491 | 2.9765 | 2.9765 | 5.8982 | 0.77072
12 | 99 | 185530.9844 | 3.4422 | 3.4422 | 5.0559 | 0.7038
14 | 117 | 208035.1499 | 4.3279 | 4.3279 | 3.9875 | 0.6007
15 | 126 | 217396.1344 | 4.7261 | 4.7261 | 4.0935 | 0.61209
16 | 134 | 224593.4137 | 5.0442 | 5.0442 | 3.6381 | 0.56088
17 | 141 | 229994.5183 | 5.2897 | 5.2897 | 3.5123 | 0.54559
18 | 147 | 233941.8381 | 5.4729 | 5.4729 | 3.0398 | 0.48285
Intercept
[10 empty] | Slope | 0.125
[10 empty] | Rg,Guinier | 61.2372 | nm
[10 empty] | R if sphere | 79.0569 | nm
[10 empty] | Rh from DLS | 92 | nm
[5 empty] | Measured | Normalized |  | CalcInten
q/cm-1 | q/nm-1 | x=qR | PthetaCalc | Icalc | q/nm-1 | intensity, I | x=qR | Izero*Icalc,pt by pt | Sq.Err.
1 | 1e-07 | 9.7207e-06 | 1 | 9.9229 | 0.00476 | 12.3272 | 0.4627 | 9.5057 | 7.9607 | Izero/Arbitrary | 9.9229
10000 | 0.001 | 0.097207 | 0.99811 | 9.9041 | 0.0059026 | 11.1703 | 0.57378 | 9.2877 | 3.5442 | Rsphere/nm | 97.2068
15000 | 0.0015 | 0.14581 | 0.99576 | 9.8808 | 0.0067253 | 10.4295 | 0.65374 | 9.1051 | 1.7538
20000 | 0.002 | 0.19441 | 0.99247 | 9.8481 | 0.00914 | 7.6296 | 0.88847 | 8.4583 | 0.6866
25000 | 0.0025 | 0.24302 | 0.98825 | 9.8063 | 0.010311 | 6.7932 | 1.0023 | 8.0925 | 1.6881
30000 | 0.003 | 0.29162 | 0.98311 | 9.7553 | 0.011642 | 8.5837 | 1.1317 | 7.6432 | 0.88444
35000 | 0.0035 | 0.34022 | 0.97708 | 9.6954 | 0.012929 | 6.4468 | 1.2568 | 7.181 | 0.53901
40000 | 0.004 | 0.38883 | 0.97015 | 9.6267 | 0.014341 | 7.2865 | 1.3941 | 6.6505 | 0.40449
45000 | 0.0045 | 0.43743 | 0.96235 | 9.5493 | 0.015846 | 5.4497 | 1.5403 | 6.0678 | 0.38215
50000 | 0.005 | 0.48603 | 0.9537 | 9.4634 | 0.017253 | 5.8982 | 1.6771 | 5.5159 | 0.14616
55000 | 0.0055 | 0.53464 | 0.94421 | 9.3693 | 0.018553 | 5.0559 | 1.8035 | 5.0068 | 0.0024137
60000 | 0.006 | 0.58324 | 0.93392 | 9.2671 | 0.020804 | 3.9875 | 2.0222 | 4.1469 | 0.025404
65000 | 0.0065 | 0.63184 | 0.92283 | 9.1572 | 0.02174 | 4.0935 | 2.1132 | 3.8026 | 0.084602
70000 | 0.007 | 0.68045 | 0.91099 | 9.0396 | 0.022459 | 3.6381 | 2.1832 | 3.545 | 0.0086681
75000 | 0.0075 | 0.72905 | 0.89841 | 8.9149 | 0.022999 | 3.5123 | 2.2357 | 3.3563 | 0.024333
80000 | 0.008 | 0.77765 | 0.88514 | 8.7831 | 0.023394 | 3.0398 | 2.2741 | 3.221 | 0.032829
85000 | 0.0085 | 0.82626 | 0.87119 | 8.6447 |  |  |  | SSE--> | 6.663
90000 | 0.009 | 0.87486 | 0.8566 | 8.4999
95000 | 0.0095 | 0.92346 | 0.8414 | 8.3491
100000 | 0.01 | 0.97207 | 0.82563 | 8.1926
105000 | 0.0105 | 1.0207 | 0.80933 | 8.0308
110000 | 0.011 | 1.0693 | 0.79252 | 7.8641
115000 | 0.0115 | 1.1179 | 0.77525 | 7.6927
120000 | 0.012 | 1.1665 | 0.75756 | 7.5172
125000 | 0.0125 | 1.2151 | 0.73948 | 7.3378
130000 | 0.013 | 1.2637 | 0.72106 | 7.155
135000 | 0.0135 | 1.3123 | 0.70233 | 6.9691
140000 | 0.014 | 1.3609 | 0.68333 | 6.7806
145000 | 0.0145 | 1.4095 | 0.6641 | 6.5897
150000 | 0.015 | 1.4581 | 0.64468 | 6.397
155000 | 0.0155 | 1.5067 | 0.6251 | 6.2028
160000 | 0.016 | 1.5553 | 0.60542 | 6.0075
165000 | 0.0165 | 1.6039 | 0.58567 | 5.8115
170000 | 0.017 | 1.6525 | 0.56587 | 5.6151
175000 | 0.0175 | 1.7011 | 0.54609 | 5.4187
180000 | 0.018 | 1.7497 | 0.52633 | 5.2227
185000 | 0.0185 | 1.7983 | 0.50666 | 5.0275
190000 | 0.019 | 1.8469 | 0.48709 | 4.8333
195000 | 0.0195 | 1.8955 | 0.46766 | 4.6405
200000 | 0.02 | 1.9441 | 0.44841 | 4.4495
205000 | 0.0205 | 1.9927 | 0.42936 | 4.2605
210000 | 0.021 | 2.0413 | 0.41055 | 4.0738
215000 | 0.0215 | 2.0899 | 0.392 | 3.8898
220000 | 0.022 | 2.1386 | 0.37375 | 3.7087
225000 | 0.0225 | 2.1872 | 0.35581 | 3.5307
230000 | 0.023 | 2.2358 | 0.33822 | 3.3561
235000 | 0.0235 | 2.2844 | 0.32099 | 3.1851
240000 | 0.024 | 2.333 | 0.30415 | 3.018
245000 | 0.0245 | 2.3816 | 0.28771 | 2.8549
250000 | 0.025 | 2.4302 | 0.27169 | 2.696

## Parameters
Solvent | Water
Wavelength/nm | 685
Wavelength/cm | 6.85e-05
Refractive index | 1.33