## Development notes
Developed May-June 2022 by Paul Russo
Made separate Guinier plots for ALV and Wyatt detectors at ORNL
Changed log(I) to ln(I)
Sep5_2023 | Modified Guinier page; cosmetic changes

## Instructions & Notes
Purpose:
This workbook permits us to normalize detectors to a standard of finite size.
Using the normalization, you can then push ahead to measure Rg using the law of Guinier
The master version contains data from a normalization and Rg
measurement (unlike a blank template). Be sure to cover these tracks with your new data.
Principle:
There is no such thing as a perfect Rayleigh scatterer. Typically, we might use
a small latex particle, bovine serum albumen, or Brij-97 microemulsion.
The idea is to use something small that scatters a lot, well above solvent level and any
stray light.
The SMALS instrument has two kinds of detectors: Wyatt pin photodiodes plus
Excelitas avalanche photodiode light sensors selected for DLS performance by ALV.
The Wyatt detectors can most easily be read from the front panel display.
More accurate readings can be obtained by collecting data using the Astra software (requires Win98).
The unit of measurement is probably in volts, but it does not matter.
The APD detectors measure photocounts in kHz, a totally different thing.
The objective of normalization is to figure out what factor puts the measured intensities for the
normalization standard on the correct P(q) curve. We calculate the curve from the low-order P(q) expansion, 1-q2Rg2/3.
That will be good enough for any sensible normalization standard, but the equation
could be replaced with a detailed expression for, say, a larger latex sphere.
TO USE:
NORMALIZATION USING A SOLUTION (E.G., BSA, BRIJ MICROEMULSION OR PS/TOL)
1 | Make a DLS run on your normalization solution, save as .ASC file.
2 | Normally, this will give you the automatic cumulants Gamma value, a Rh value, and intensities for each APD detector. Copy the intensities.
3 | Select the NormStandard sheet; paste the intensities from .ASC file into the yellow cells (cursor on Q16). The ALV results are automatically copied where they need to go.
4 | Read the ALV's front panel display for the laser monitor and all angles by gently pressing the selection button;
type the results into the blue cells in column H.
5 | Repeat the previous 2 steps for the normalization standard's solvent (usually a much weaker scatterer) blue cells of column K.
but paste intensites to cell T16 and type the panel readings into
6 | Enter the Rh value (from .ASC file's automatic cumulant zone or any other analysis) into the K7 cell.
7 | Enter your name and date in the appropriate green cells.
8 | Enter "Y' or "N" in F6 to indicate whether or not beamstop was used.
9 | Enter the name of normalization standard in F3.
10 | At this point, your normalization constants should be ready.
SIZE MEASUREMENT OF A DISSOLVED SOLUTE
1 | Select the Conc1 sheet & measure your solution and solvent. For the ALV APD readings, enter intensities in cells O16 or R16, respectively.
2 | For the Wyatt diode detectors enter the front panel readings in the blue cells of columns G and J, respectively.
3 | You can enter the solute's Rh value from DLS in cell J7; a DLS-based Rg estimate will be calculated automatically from a sphere model,
along with comparisons to wavelength and the P(q) estimate at max angle; these will tell you whether there is hope of getting a useful size.
4 | A Guinier plot or similar should give you the SLS size.
5 | Then you can get Rg/Rh for shape analysis (e.g., should be sqrt(3/5) for solid, uniform spheres).
REPEAT FOR EACH CONCENTRATION
EXTRAPOLATE THE SIZE RESULTS TO c = 0
Caveat: | Similar to a GUI program, it is incumbent on the user to remember to enter data, update labels, etc.
In a good GUI program (e.g., Visual C) it is easy to flag users when data is not updated or omitted.
That's a little trickier in VB for Excel, but probably doable. It's just not done here, that's all.
Limitation: | In this Master document, the data are populated only for the Excelitas APD detectors.
Zimm plots: | It has been many years since the Zimm plot was an efficient way to process SLS data, but
it's still a great way to present a snapshot of a polymer's molecular weight, size and virial coefficient.
This assertion is developed in our J. Chem. Ed. Article: (see: https://pubs.acs.org/doi/abs/10.1021/ed076p1534 ).
Data from NormaSize can be extracted for Zimm plot creation as follows.
1. Perform normalization as usual (or inherit a previous result by making a new copy of NormaSize)
2. Measure the sample and solvent as described above and briefly repeated below.
3. Data from the Excelitas APDs can be copy/pasted from the .ASC file into the marked boxes -->
That will automatically populate the appropriate cells for those detectors
4. Data from the Wyatt pin photodiodes can be entered manually. At low angles,
it may help to average several readings (in teaching situations, several people
can provide a reading).
It is also possible to use Wyatt Astra software to record these values. Refer to the Astra manual.
5. The angle/intensity data are in columns D and M, respectively.
6
Rayleigh factors: |  | The equations and figure below come from the Macro HowTo: https://macro.lsu.edu/HowTo/guide/guide.doc
These equations are intended for rotating detectors.
We will modify them the last one for SMALS.
An important caveat here about using the SMALS apparatus: we have only
listed the Rayleigh factors at 90 degrees scattering angle. There are no
APD detectors at this angle. So, even if the plan is to use just the APD detectors
for SLS, one must still make a measurement at 90 degrees with the
Wyatt's diode detector at that angle.

## Parameters
Solvent for solutions | Toluene |  | Rayleigh calibration standard | Toluene | Almost always toluene | Traditional values we have always used (e.g., GuiDe documentation) are scaled by lo-4 from a 632.8 nm result, probably Kaye & McDaniel.
Wavelength/nm | 685 |  | Rayleigh factor for calib. Std. | 9.9584e-05 | cm-1; reminder: | Rtol,Uv = | 14.02 x 10-6 cm-1 @ 632.8 nm (HeNe red)
Wavelength/cm | 6.85e-05 |  | Measured intensity Rayleigh Std at 90 deg. | 0.03 | (from Wyatt front panel display) |  | 28.07 x 10-6 cm-1 @ 532 nm (Frequency doubled diode)
Refractive index of solutions & solvent | 1.495 |  | Volume exponent, m | 1 |  |  | 32.08 x 10-6 cm-1 @ 514.5 nm (Argon ion green)
Refactive index of Rayleigh standard | 1.495 | [5 empty] | 39.64 x 10-6 cm-1 @ 488 nm (Argon ion blue)
Volume expansion exponent, m | 1
Rayleigh factor of toluene | 1 | << | See values in N11 - N15 for common wavelengths |  |  | These traditional values assume a lo-4 dependence. The wavelength dependence was re-examined to account for
dn/dc in mL/gram | 0.11 |  |  |  |  | refractive index dispersion by Wu* and the values claimed for toluene are:
Optical constant K/cm2.mol.g-2 | 8.0524e-08 |  |  |  |  | Uu geometry* |  |  |  | Uv geometry (** our own calculation, based on Wu's collection of literature data, adjusted from Uu to Uv)
[6 empty] | Rtol,Uu = 4.9 x 106 /lo4.17 * | in cm-1 with lo given in nm |  |  | Rtol,Uv = 6.653 x 106 /lo4.17 ** |  | in cm-1 with lo in nm
[6 empty] | lo/nm | R in cm-1 |  |  | These values were calculated using |  |  |  |  | << use lo not l in all these expressions
[6 empty] | 685 | 7.3344e-06 | cm-1 |  | 9.9584e-05 | A0 | A1 | A2 | A3 | A4
[6 empty] | 632.8 | 1.0207e-05 |  |  | 0.00013859 | 2.1617 | 0.00049519 | 0.021382 | 5.8838e-05 | 8.7632e-05
[6 empty] | 532 | 2.1045e-05 |  |  | 0.00028574 | assuming ru = 0.495 from the data collected by Hu
[6 empty] | 514.5 | 2.4195e-05 |  |  | 0.0003285 | The dispersion equation is from: |  |  | Moutzouris, K., Papamichael, M., Betsis, S.C. et al. Refractive, dispersive and thermo-optic properties of twelve organic solvents in the visible and near-infrared. Appl. Phys. B 116, 617–622 (2014). https://doi.org/10.1007/s00340-013-5744-3
[6 empty] | 488 | 3.0164e-05 |  |  | 0.00040955 | See our spreadsheet WavelengthDispersion.xlsx and associated Origin files RayleighVsWavelength.opj (2 files with different error assumptions)
* H. Wu, Chemical Physics 367 (1), 44-47 (2010), https://doi.org/10.1016/j.chemphys.2009.10.019
** RUv analysis from our own WavelengthDispersion spreadsheet & associated Origin non-linear fit to data with error bars
The 4.17 exponent values in N11 - N15 compare reasonably well to the 4.0 exponent values in K2 - K5
Should they? Yes...here is how a 4.0 and 4.17 exponent differ.
[9 empty] | (633/488)^4 = | 2.831
[9 empty] | (633/488)^4.17 = | 2.959
[9 empty] | % Difference | 4.3263 | %
From JustWyatt xls spreadsheet April 6, 2024
Note on B7: Author: Not really sure whether the exponent should be 1 or 2

## ConcData
SAMPLE PREPARATION BY WEIGHT (PARTIAL SPECIFIC VOLUME OF SOLUTE & SOLVENT DENSITY)
Solute: | PS 170 pellets
Solvent: | Toluene
Solvent filter size and type: | PALL 0.2 um nylon filters
Partial specific volume v2bar | 0.918 | mL/g
partial specific vol reference | https://nvlpubs.nist.gov/nistpubs/jres/52/jresv52n4p217_a1b.pdf
solvent density | 0.86 | mL/g
solvent reference | https://nvlpubs.nist.gov/nistpubs/jres/52/jresv52n4p217_a1b.pdf
measurement temperature | 25 | oC
|  | sample label | c1 | c2 | c3 | c4 | c5 | c6 | c7 | c8
|  | pellets, if applicable | 1 | 2 | 3 | 4 | 0 | 0 | 0 | 0
|  | est. tol, mL | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
|  | scint vial | 12.9162 | 12.8436 | 13.0953 | 13.0998 | 0 | 0 | 0 | 0
|  | plus pellet | 12.9492 | 12.9029 | 13.1848 | 13.2256 | 0 | 0 | 0 | 0
|  | plus toluene | 23.8342 | 23.4196 | 22.5286 | 23.6966 | 0 | 0 | 0 | 0
|  | solute | 0.033 | 0.0593 | 0.0895 | 0.1258 | 0 | 0 | 0 | 0
|  | solvent | 10.885 | 10.5167 | 9.3438 | 10.471 | 0 | 0 | 0 | 0
|  | c/g·mL-1 | 0.002601 | 0.0048277 | 0.0081757 | 0.010235 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
notes & observations |  | I measured at 25C, but the v2bar and rho1 data are at ~27C
For c2 and c3, probably the filter was broken....catch the dust later
adjustments (solvent evaporation, centrifugation or filtration to remove dust/aggregates) | bar: c1 to c3
None
|  | cfinal/g·mL-1 | 0.002601 | 0.0048277 | 0.0081757 | 0.010235 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## NormStandard
Date of normalization |  |  | March 2 2023
Normalizing standard |  |  | Brij
Operator |  |  | Russo
Comment¯ | Monitors
| Front |  | beam stop (y/n) | y | 0.015 |  | Parameter
| p1 |  |  |  | 4.98 |  | Rh | 4.76 | nm
| p2 |  |  |  | 14.99 |  | Rg (sphere) | 3.6871 | nm
| p3 |  |  |  | -14.8 |  | lo/(n*DLS radius) | 96.2592
| LC |  |  |  | 79.4 |  | qmaxRg | 0.096957
| L |  |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99687
Normalization
Comment¯ | Port # | SMALS Det # | Angle/degrees |  | ---- | Normalizing solution |  | Normalizing solvent background |  | Signal | P(q) | Factor
| Zero-ref |  | 0 |  | Detector type | Wyatt Reading | ALV Reading/kHz | Wyatt Reading | ALV Reading |  | 1 | 1 | Drop solution count rates |  | Drop solvent count rates
| 1 | 1 | 22.5 |  | Wyatt diode | 1.0337 |  | 0.02 |  | 1.0137 | 0.99987 | 0.98631 | here from .ASC file |  | here from .ASC file |  | April 5, 2024 Brij |  | April 5, 2024 Dirty Water
| 2 | 2 | 28 |  | Wyatt diode | 0.69875 |  | 0.06 |  | 0.63875 | 0.9998 | 1.5652 | MeanCR1 [kHz] : | 261.941 | MeanCR1 [kHz] : | 1.9109 | MeanCR1 [kHz] : | 489.1353 | MeanCR1 [kHz] : | 2.2987
| 3 | 3 | 32 |  | Wyatt diode | 0.85 |  | 0.015 |  | 0.835 | 0.99973 | 1.1973 | MeanCR2 [kHz] : | 872.9901 | MeanCR2 [kHz] : | 3.1774 | MeanCR2 [kHz] : | 793.152 | MeanCR2 [kHz] : | 3.5894
| 4 |  | 38 |  | Empty | --- |  | --- |  | --- | --- | --- | MeanCR3 [kHz] : | 796.7864 | MeanCR3 [kHz] : | 3.1847 | MeanCR3 [kHz] : | 746.4744 | MeanCR3 [kHz] : | 1.6484
| 5 | 4 | 44 |  | Fiber/APD |  | 261.941 |  | 1.9109 | 260.0301 | 0.9995 | 0.0038438 | MeanCR4 [kHz] : | 819.8116 | MeanCR4 [kHz] : | 10.3614 | MeanCR4 [kHz] : | 666.2175 | MeanCR4 [kHz] : | 13.6203
| 6 | 5 | 50 |  | Fiber/APD |  | 872.9901 |  | 3.1774 | 869.8127 | 0.99935 | 0.0011489 | MeanCR5 [kHz] : | 352.5294 | MeanCR5 [kHz] : | 1.2127 | MeanCR5 [kHz] : | 188.6304 | MeanCR5 [kHz] : | 1.2444
| 7 | 6 | 57 |  | Wyatt diode | 1.161 |  | 0.027 |  | 1.134 | 0.99916 | 0.88109 | MeanCR6 [kHz] : | 328.7788 | MeanCR6 [kHz] : | 1.228 | MeanCR6 [kHz] : | 322.4396 | MeanCR6 [kHz] : | 1.5229
| 8 | 7 | 64 |  | Fiber/APD |  | 796.7864 |  | 3.1847 | 793.6017 | 0.99894 | 0.0012587 | MeanCR7 [kHz] : | 311.9114 | MeanCR7 [kHz] : | 1.7291 | MeanCR7 [kHz] : | 497.8206 | MeanCR7 [kHz] : | 2.6058
| 9 | 9 | 72 |  | Wyatt diode | 1.056 |  | 0.009 |  | 1.047 | 0.99865 | 0.95382 | MeanCR8 [kHz] : | 572.8061 | MeanCR8 [kHz] : | 1.8599 | MeanCR8 [kHz] : | 557.6425 | MeanCR8 [kHz] : | 2.7941
| 10 | 9 | 81 |  | Fiber/APD |  | 819.8116 |  | 10.3614 | 809.4503 | 0.9983 | 0.0012333
| 11 | 10 | 90 |  | Wyatt diode | 0.948 |  | 0.007 |  | 0.941 | 0.9979 | 1.0605
| 12 | 11 | 99 |  | Fiber/APD |  | 352.5294 |  | 1.2127 | 351.3167 | 0.99746 | 0.0028392
| 13 |  | 108 |  | Empty | --- |  | --- |  | --- | --- | ---
| 14 | 12 | 117 |  | Fiber/APD |  | 328.7788 |  | 1.228 | 327.5508 | 0.99645 | 0.0030421
| 15 | 13 | 126 |  | Wyatt diode | 1.216 |  | 0.017 |  | 1.199 | 0.99588 | 0.83059
| 16 | 14 | 134 |  | Fiber/APD |  | 311.9114 |  | 1.7291 | 310.1822 | 0.99534 | 0.0032089
| 17 | 15 | 141 |  | Wyatt diode | 0.667 |  | 0.012 |  | 0.655 | 0.99484 | 1.5188
| 18 | 16 | 147 |  | Fiber/APD |  | 572.8061 |  | 1.8599 | 570.9462 | 0.99439 | 0.0017417
Probably from an
old run.
Wyatt entries possibly from an old run.

## Conc1
Date of measurement |  | 2024-04-05
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Will SLS work?
| p1 |  |  | 4.98 |  | Rh (if avail from DLS) | 1 | nm
| p2 |  |  | 15.04 |  | Rg estimate (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmax*Estimated Rg | 0.020369
| L |  |  | 4.92 |  | 1-(qmax*estimated Rg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Drop solvent count rates
| 1 | 22.5 |  | Wyatt diode | 0.147 |  | 0.086 |  | 0.061 | 0.060165 | here from .ASC file |  | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.087 |  | 0.049 |  | 0.038 | 0.059479 | MeanCR1 [kHz] : | 30.4352 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode | 0.104 |  | 0.03 |  | 0.074 | 0.088599 | MeanCR2 [kHz] : | 93.7438 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 82.2466 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 30.4352 |  | 14.1403 | 16.2949 | 0.062634 | MeanCR4 [kHz] : | 96.0828 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 93.7438 |  | 20.9973 | 72.7465 | 0.08358 | MeanCR5 [kHz] : | 36.8912 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode | 0.13 |  | 0.051 |  | 0.079 | 0.069606 | MeanCR6 [kHz] : | 37.8758 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 82.2466 |  | 25.0557 | 57.1909 | 0.071988 | MeanCR7 [kHz] : | 32.1529 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode | 0.116 |  | 0.034 |  | 0.082 | 0.078214 | MeanCR8 [kHz] : | 67.0287 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 96.0828 |  | 43.3106 | 52.7722 | 0.065084
| 11 | 90 |  | Wyatt diode | 0.101 |  | 0.03 |  | 0.071 | 0.075293
| 12 | 99 |  | Fiber/APD |  | 36.8912 |  | 3.1831 | 33.7081 | 0.095704
| 13 | 108 |  | Empty |  |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 37.8758 |  | 5.2177 | 32.6582 | 0.09935
| 15 | 126 |  | Wyatt diode | 0.127 |  | 0.049 |  | 0.078 | 0.064786
| 16 | 134 |  | Fiber/APD |  | 32.1529 |  | 9.1797 | 22.9732 | 0.073718
| 17 | 141 |  | Wyatt diode | 0.07 |  | 0.025 |  | 0.045 | 0.068348
| 18 | 147 |  | Fiber/APD |  | 67.0287 |  | 13.5941 | 53.4346 | 0.093065
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure. Roughly speaking, anything lower than 0.9 is a good sign.

## Conc2
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode | 0.261 |  | 0.086 |  | 0.175 | 0.1726 | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.115 |  | 0.049 |  | 0.066 | 0.10331 | MeanCR1 [kHz] : | 60.7814 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode | 0.135 |  | 0.03 |  | 0.105 | 0.12572 | MeanCR2 [kHz] : | 102.0077 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty |  |  | --- |  | --- | --- | MeanCR3 [kHz] : | 87.729 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 60.7814 |  | 14.1403 | 46.641 | 0.17928 | MeanCR4 [kHz] : | 89.4668 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 102.0077 |  | 20.9973 | 81.0104 | 0.093075 | MeanCR5 [kHz] : | 21.8386 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode | 0.162 |  | 0.051 |  | 0.111 | 0.097801 | MeanCR6 [kHz] : | 39.1134 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 87.729 |  | 25.0557 | 62.6734 | 0.078889 | MeanCR7 [kHz] : | 56.0125 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode | 0.142 |  | 0.034 |  | 0.108 | 0.10301 | MeanCR8 [kHz] : | 62.7523 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 89.4668 |  | 43.3106 | 46.1563 | 0.056925
| 11 | 90 |  | Wyatt diode | 0.125 |  | 0.03 |  | 0.095 | 0.10074
| 12 | 99 |  | Fiber/APD |  | 21.8386 |  | 3.1831 | 18.6555 | 0.052966
| 13 | 108 |  | Empty |  |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 39.1134 |  | 5.2177 | 33.8957 | 0.10311
| 15 | 126 |  | Wyatt diode | 0.159 |  | 0.049 |  | 0.11 | 0.091365
| 16 | 134 |  | Fiber/APD |  | 56.0125 |  | 9.1797 | 46.8328 | 0.15028
| 17 | 141 |  | Wyatt diode | 0.087 |  | 0.025 |  | 0.062 | 0.094168
| 18 | 147 |  | Fiber/APD |  | 62.7523 |  | 13.5941 | 49.1582 | 0.085617
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Conc3
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode | 0.467 |  | 0.086 |  | 0.381 | 0.37578 | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.18 |  | 0.049 |  | 0.131 | 0.20505 | MeanCR1 [kHz] : | 77.8755 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode | 0.204 |  | 0.03 |  | 0.174 | 0.20833 | MeanCR2 [kHz] : | 126.001 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty |  |  | --- |  | --- | --- | MeanCR3 [kHz] : | 101.2723 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 77.8755 |  | 14.1403 | 63.7351 | 0.24498 | MeanCR4 [kHz] : | 94.9891 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 126.001 |  | 20.9973 | 105.0038 | 0.12064 | MeanCR5 [kHz] : | 22.5883 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode | 0.196 |  | 0.051 |  | 0.145 | 0.12776 | MeanCR6 [kHz] : | 42.7686 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 101.2723 |  | 25.0557 | 76.2166 | 0.095937 | MeanCR7 [kHz] : | 61.2646 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode | 0.165 |  | 0.034 |  | 0.131 | 0.12495 | MeanCR8 [kHz] : | 71.3666 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 94.9891 |  | 43.3106 | 51.6786 | 0.063735
| 11 | 90 |  | Wyatt diode | 0.147 |  | 0.03 |  | 0.117 | 0.12407
| 12 | 99 |  | Fiber/APD |  | 22.5883 |  | 3.1831 | 19.4052 | 0.055095
| 13 | 108 |  | Empty |  |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 42.7686 |  | 5.2177 | 37.5509 | 0.11423
| 15 | 126 |  | Wyatt diode | 0.186 |  | 0.049 |  | 0.137 | 0.11379
| 16 | 134 |  | Fiber/APD |  | 61.2646 |  | 9.1797 | 52.0849 | 0.16713
| 17 | 141 |  | Wyatt diode | 0.104 |  | 0.025 |  | 0.079 | 0.11999
| 18 | 147 |  | Fiber/APD |  | 71.3666 |  | 13.5941 | 57.7725 | 0.10062
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Conc4
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode | 0.597 |  | 0.086 |  | 0.511 | 0.504 | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.27 |  | 0.049 |  | 0.221 | 0.34592 | MeanCR1 [kHz] : | 79.0093 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode | 0.24 |  | 0.03 |  | 0.21 | 0.25143 | MeanCR2 [kHz] : | 123.1549 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty |  |  | --- |  | --- | --- | MeanCR3 [kHz] : | 111.1135 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 79.0093 |  | 14.1403 | 64.869 | 0.24934 | MeanCR4 [kHz] : | 103.1954 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 123.1549 |  | 20.9973 | 102.1576 | 0.11737 | MeanCR5 [kHz] : | 24.551 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode | 0.227 |  | 0.051 |  | 0.176 | 0.15507 | MeanCR6 [kHz] : | 45.6264 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 111.1135 |  | 25.0557 | 86.0578 | 0.10832 | MeanCR7 [kHz] : | 67.1985 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode | 0.179 |  | 0.034 |  | 0.145 | 0.1383 | MeanCR8 [kHz] : | 80.5785 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 103.1954 |  | 43.3106 | 59.8848 | 0.073856
| 11 | 90 |  | Wyatt diode | 0.158 |  | 0.03 |  | 0.128 | 0.13574
| 12 | 99 |  | Fiber/APD |  | 24.551 |  | 3.1831 | 21.3679 | 0.060668
| 13 | 108 |  | Empty |  |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 45.6264 |  | 5.2177 | 40.4087 | 0.12293
| 15 | 126 |  | Wyatt diode | 0.208 |  | 0.049 |  | 0.159 | 0.13206
| 16 | 134 |  | Fiber/APD |  | 67.1985 |  | 9.1797 | 58.0188 | 0.18618
| 17 | 141 |  | Wyatt diode | 0.118 |  | 0.025 |  | 0.093 | 0.14125
| 18 | 147 |  | Fiber/APD |  | 80.5785 |  | 13.5941 | 66.9844 | 0.11666
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Conc5
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode |  |  | 0.086 |  | -0.086 | -0.084822 | here from .ASC file
| 2 | 28 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.076697 | MeanCR1 [kHz] : | 30.4352 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.035919 | MeanCR2 [kHz] : | 93.7438 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 82.2466 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 30.4352 |  | 14.1403 | 16.2949 | 0.062634 | MeanCR4 [kHz] : | 96.0828 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 93.7438 |  | 20.9973 | 72.7465 | 0.08358 | MeanCR5 [kHz] : | 36.8912 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode |  |  | 0.051 |  | -0.051 | -0.044936 | MeanCR6 [kHz] : | 37.8758 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 82.2466 |  | 25.0557 | 57.1909 | 0.071988 | MeanCR7 [kHz] : | 32.1529 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode |  |  | 0.034 |  | -0.034 | -0.03243 | MeanCR8 [kHz] : | 67.0287 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 96.0828 |  | 43.3106 | 52.7722 | 0.065084
| 11 | 90 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.031814
| 12 | 99 |  | Fiber/APD |  | 36.8912 |  | 3.1831 | 33.7081 | 0.095704
| 13 | 108 |  | Empty | --- |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 37.8758 |  | 5.2177 | 32.6582 | 0.09935
| 15 | 126 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.040699
| 16 | 134 |  | Fiber/APD |  | 32.1529 |  | 9.1797 | 22.9732 | 0.073718
| 17 | 141 |  | Wyatt diode |  |  | 0.025 |  | -0.025 | -0.037971
| 18 | 147 |  | Fiber/APD |  | 67.0287 |  | 13.5941 | 53.4346 | 0.093065
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Conc6
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode |  |  | 0.086 |  | -0.086 | -0.084822 | here from .ASC file
| 2 | 28 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.076697 | MeanCR1 [kHz] : | 30.4352 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.035919 | MeanCR2 [kHz] : | 93.7438 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 82.2466 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 30.4352 |  | 14.1403 | 16.2949 | 0.062634 | MeanCR4 [kHz] : | 96.0828 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 93.7438 |  | 20.9973 | 72.7465 | 0.08358 | MeanCR5 [kHz] : | 36.8912 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode |  |  | 0.051 |  | -0.051 | -0.044936 | MeanCR6 [kHz] : | 37.8758 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 82.2466 |  | 25.0557 | 57.1909 | 0.071988 | MeanCR7 [kHz] : | 32.1529 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode |  |  | 0.034 |  | -0.034 | -0.03243 | MeanCR8 [kHz] : | 67.0287 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 96.0828 |  | 43.3106 | 52.7722 | 0.065084
| 11 | 90 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.031814
| 12 | 99 |  | Fiber/APD |  | 36.8912 |  | 3.1831 | 33.7081 | 0.095704
| 13 | 108 |  | Empty | --- |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 37.8758 |  | 5.2177 | 32.6582 | 0.09935
| 15 | 126 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.040699
| 16 | 134 |  | Fiber/APD |  | 32.1529 |  | 9.1797 | 22.9732 | 0.073718
| 17 | 141 |  | Wyatt diode |  |  | 0.025 |  | -0.025 | -0.037971
| 18 | 147 |  | Fiber/APD |  | 67.0287 |  | 13.5941 | 53.4346 | 0.093065
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Conc7
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode |  |  | 0.086 |  | -0.086 | -0.084822 | here from .ASC file
| 2 | 28 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.076697 | MeanCR1 [kHz] : | 30.4352 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.035919 | MeanCR2 [kHz] : | 93.7438 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 82.2466 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 30.4352 |  | 14.1403 | 16.2949 | 0.062634 | MeanCR4 [kHz] : | 96.0828 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 93.7438 |  | 20.9973 | 72.7465 | 0.08358 | MeanCR5 [kHz] : | 36.8912 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode |  |  | 0.051 |  | -0.051 | -0.044936 | MeanCR6 [kHz] : | 37.8758 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 82.2466 |  | 25.0557 | 57.1909 | 0.071988 | MeanCR7 [kHz] : | 32.1529 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode |  |  | 0.034 |  | -0.034 | -0.03243 | MeanCR8 [kHz] : | 67.0287 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 96.0828 |  | 43.3106 | 52.7722 | 0.065084
| 11 | 90 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.031814
| 12 | 99 |  | Fiber/APD |  | 36.8912 |  | 3.1831 | 33.7081 | 0.095704
| 13 | 108 |  | Empty | --- |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 37.8758 |  | 5.2177 | 32.6582 | 0.09935
| 15 | 126 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.040699
| 16 | 134 |  | Fiber/APD |  | 32.1529 |  | 9.1797 | 22.9732 | 0.073718
| 17 | 141 |  | Wyatt diode |  |  | 0.025 |  | -0.025 | -0.037971
| 18 | 147 |  | Fiber/APD |  | 67.0287 |  | 13.5941 | 53.4346 | 0.093065
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Conc8
Date of measurement |  | 2024-04-05 |  |  | Metadata and solvent intensity values are inherited from the Conc1 worksheet
Sample |  | PS Cmax 100
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | Y | 0.006 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 1 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 0.7746 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 458.194
| LC |  |  | 79.2 |  | qmaxRg | 0.020369
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99986
Normalized intensity
Comment¯ | Port # | Angle/degrees |  | ---- | Solution reading |  | Solvent background |  | Signal
| Zero-ref | 0 |  | Detector type | Wyatt reading | ALV reading/kHz | Wyatt reading | ALV reading |  |  | Drop solution count rates |  | Taken from Conc1 Sheet
| 1 | 22.5 |  | Wyatt diode |  |  | 0.086 |  | -0.086 | -0.084822 | here from .ASC file
| 2 | 28 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.076697 | MeanCR1 [kHz] : | 30.4352 | MeanCR1 [kHz] : | 14.1403
| 3 | 32 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.035919 | MeanCR2 [kHz] : | 93.7438 | MeanCR2 [kHz] : | 20.9973
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 82.2466 | MeanCR3 [kHz] : | 25.0557
| 5 | 44 |  | Fiber/APD |  | 30.4352 |  | 14.1403 | 16.2949 | 0.062634 | MeanCR4 [kHz] : | 96.0828 | MeanCR4 [kHz] : | 43.3106
| 6 | 50 |  | Fiber/APD |  | 93.7438 |  | 20.9973 | 72.7465 | 0.08358 | MeanCR5 [kHz] : | 36.8912 | MeanCR5 [kHz] : | 3.1831
| 7 | 57 |  | Wyatt diode |  |  | 0.051 |  | -0.051 | -0.044936 | MeanCR6 [kHz] : | 37.8758 | MeanCR6 [kHz] : | 5.2177
| 8 | 64 |  | Fiber/APD |  | 82.2466 |  | 25.0557 | 57.1909 | 0.071988 | MeanCR7 [kHz] : | 32.1529 | MeanCR7 [kHz] : | 9.1797
| 9 | 72 |  | Wyatt diode |  |  | 0.034 |  | -0.034 | -0.03243 | MeanCR8 [kHz] : | 67.0287 | MeanCR8 [kHz] : | 13.5941
| 10 | 81 |  | Fiber/APD |  | 96.0828 |  | 43.3106 | 52.7722 | 0.065084
| 11 | 90 |  | Wyatt diode |  |  | 0.03 |  | -0.03 | -0.031814
| 12 | 99 |  | Fiber/APD |  | 36.8912 |  | 3.1831 | 33.7081 | 0.095704
| 13 | 108 |  | Empty | --- |  | --- | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 37.8758 |  | 5.2177 | 32.6582 | 0.09935
| 15 | 126 |  | Wyatt diode |  |  | 0.049 |  | -0.049 | -0.040699
| 16 | 134 |  | Fiber/APD |  | 32.1529 |  | 9.1797 | 22.9732 | 0.073718
| 17 | 141 |  | Wyatt diode |  |  | 0.025 |  | -0.025 | -0.037971
| 18 | 147 |  | Fiber/APD |  | 67.0287 |  | 13.5941 | 53.4346 | 0.093065
Note on J7: Author: From DLS results; enter 1 if not yet known
Note on J9: Author: Light wavelength relative to scatterer size
Note on J10: Author: <1 implies you are in Guinier region
Note on J11: Author: The deviation from 1 shows how much there is to measure.

## Guinier+SphereFit1
Color key | Wyatt Diode
ALV APD
| Port | Angle/degrees | q/cm-1 | q2/1010cm-2 | q2/10-4nm-2 | Normalized Intensity, I | ln (I)
| 1 | 22.5 | 53505.227 | 0.28628 | 0.28628 | 0.060165 | -2.8107
| 2 | 28 | 66349.1956 | 0.44022 | 0.44022 | 0.059479 | -2.8221
| 3 | 32 | 75595.9554 | 0.57147 | 0.57147 | 0.088599 | -2.4236
| 5 | 44 | 102739.1344 | 1.0555 | 1.0555 | 0.062634 | -2.7704 | [7 empty] | ALV APD |  |  |  |  | Normalized |  |  |  | Wyatt diode |  |  |  |  | Normalized
| 6 | 50 | 115906.7543 | 1.3434 | 1.3434 | 0.08358 | -2.4819 | [6 empty] | Row # | Port | q/degrees | q/cm-1 | q2/1010cm-2 | q2/10-4nm-2 | intensity, I | ln (I) |  | Row # | Port | Angle/degrees | q/cm-1 | q2/1010cm-2 | q2/10-4nm-2 | intensity, I | ln (I)
| 7 | 57 | 130864.963 | 1.7126 | 1.7126 | 0.069606 | -2.6649 | [6 empty] | 10 | 5 | 44 | 102739.1344 | 1.0555 | 1.0555 | 0.062634 | -2.7704 |  | 7 | 1 | 22.5 | 53505.227 | 0.28628 | 0.28628 | 0.060165 | -2.8107
| 8 | 64 | 145334.9926 | 2.1122 | 2.1122 | 0.071988 | -2.6313 | [6 empty] | 11 | 6 | 50 | 115906.7543 | 1.3434 | 1.3434 | 0.08358 | -2.4819 |  | 8 | 2 | 28 | 66349.1956 | 0.44022 | 0.44022 | 0.059479 | -2.8221
| 9 | 72 | 161205.2459 | 2.5987 | 2.5987 | 0.078214 | -2.5483 | [6 empty] | 13 | 8 | 64 | 145334.9926 | 2.1122 | 2.1122 | 0.071988 | -2.6313 |  | 9 | 3 | 32 | 75595.9554 | 0.57147 | 0.57147 | 0.088599 | -2.4236
| 10 | 81 | 178116.807 | 3.1726 | 3.1726 | 0.065084 | -2.7321 | [6 empty] | 15 | 10 | 81 | 178116.807 | 3.1726 | 3.1726 | 0.065084 | -2.7321 |  | 12 | 7 | 57 | 130864.963 | 1.7126 | 1.7126 | 0.069606 | -2.6649
| 11 | 90 | 193930.2188 | 3.7609 | 3.7609 | 0.075293 | -2.5864 | [6 empty] | 17 | 12 | 99 | 208547.9862 | 4.3492 | 4.3492 | 0.095704 | -2.3465 |  | 14 | 9 | 72 | 161205.2459 | 2.5987 | 2.5987 | 0.078214 | -2.5483
| 12 | 99 | 208547.9862 | 4.3492 | 4.3492 | 0.095704 | -2.3465 | [6 empty] | 18 | 14 | 117 | 233844.0219 | 5.4683 | 5.4683 | 0.09935 | -2.3091 |  | 16 | 11 | 90 | 193930.2188 | 3.7609 | 3.7609 | 0.075293 | -2.5864
| 14 | 117 | 233844.0219 | 5.4683 | 5.4683 | 0.09935 | -2.3091 | [6 empty] | 20 | 16 | 134 | 252456.5064 | 6.3734 | 6.3734 | 0.073718 | -2.6075 |  | 19 | 15 | 126 | 244366.3316 | 5.9715 | 5.9715 | 0.064786 | -2.7367
| 15 | 126 | 244366.3316 | 5.9715 | 5.9715 | 0.064786 | -2.7367 | [6 empty] | 22 | 18 | 147 | 262964.6977 | 6.915 | 6.915 | 0.093065 | -2.3745 |  | 21 | 17 | 141 | 258527.6728 | 6.6837 | 6.6837 | 0.068348 | -2.6831
| 16 | 134 | 252456.5064 | 6.3734 | 6.3734 | 0.073718 | -2.6075
| 17 | 141 | 258527.6728 | 6.6837 | 6.6837 | 0.068348 | -2.6831
| 18 | 147 | 262964.6977 | 6.915 | 6.915 | 0.093065 | -2.3745 |  |  |  | All Detectors Results
[11 empty] | Intercept | -2.6774 | arb. | [6 empty] | ALV APD Results | [6 empty] | Wyatt diode results
[11 empty] | -Slope | -0.024787 | nm2 | [6 empty] | Intercept | -2.7012 | arb. |  |  |  |  | Intercept | -2.6567 | arb.
[11 empty] | Rg,Guinier | 27.2692 | nm | [6 empty] | -Slope | -0.044047 | nm2 |  |  |  |  | -Slope | -0.001012 | nm2
[11 empty] | R if sphere | 35.2044 | nm | [6 empty] | Rg,Guinier | 36.3511 | nm |  |  |  |  | Rg,Guinier | 5.51 | nm
[11 empty] | Rh from DLS | 1 | nm | [6 empty] | R if sphere | 46.9291 | nm |  |  |  |  | R if sphere | 7.1134 | nm
[6 empty] | Measured | Normalized |  | CalcInten | [10 empty] | Rh from DLS | 1 | nm |  |  |  |  | Rh from DLS | 1 | nm
| q/cm-1 | q/nm-1 | x=qR | PthetaCalc | Icalc | q/nm-1 | intensity, I | x=qR | Izero*Icalc,pt by pt | Sq.Err.
| 1 | 1e-07 | 1.1269e-05 | 1 | 1.4457 | 0.0053505 | 0.060165 | 0.60295 | 1.3438 | 1.6477 | Izero/Arbitrary | 1.4457
| 10000 | 0.001 | 0.11269 | 0.99746 | 1.442 | 0.0066349 | 0.059479 | 0.74769 | 1.2916 | 1.5181 | Rsphere/nm | 112.6903
| 15000 | 0.0015 | 0.16904 | 0.9943 | 1.4374 | 0.0075596 | 0.088599 | 0.85189 | 1.2484 | 1.3452
| 20000 | 0.002 | 0.22538 | 0.98988 | 1.4311 | 0.010274 | 0.062634 | 1.1578 | 1.0998 | 1.0757
| 25000 | 0.0025 | 0.28173 | 0.98423 | 1.4229 | 0.011591 | 0.08358 | 1.3062 | 1.0188 | 0.87459
| 30000 | 0.003 | 0.33807 | 0.97736 | 1.413 | 0.013086 | 0.069606 | 1.4747 | 0.92234 | 0.72715
| 35000 | 0.0035 | 0.39442 | 0.9693 | 1.4013 | 0.014533 | 0.071988 | 1.6378 | 0.82674 | 0.56966
| 40000 | 0.004 | 0.45076 | 0.96006 | 1.3879 | 0.016121 | 0.078214 | 1.8166 | 0.72179 | 0.41419
| 45000 | 0.0045 | 0.50711 | 0.94969 | 1.3729 | 0.017812 | 0.065084 | 2.0072 | 0.61259 | 0.29976
| 50000 | 0.005 | 0.56345 | 0.93821 | 1.3563 | 0.019393 | 0.075293 | 2.1854 | 0.51531 | 0.19362
| 55000 | 0.0055 | 0.6198 | 0.92565 | 1.3382 | 0.020855 | 0.095704 | 2.3501 | 0.43123 | 0.11258
| 60000 | 0.006 | 0.67614 | 0.91207 | 1.3186 | 0.023384 | 0.09935 | 2.6352 | 0.30234 | 0.041203
| 65000 | 0.0065 | 0.73249 | 0.8975 | 1.2975 | 0.024437 | 0.064786 | 2.7538 | 0.25569 | 0.036446
| 70000 | 0.007 | 0.78883 | 0.88199 | 1.2751 | 0.025246 | 0.073718 | 2.8449 | 0.22278 | 0.022218
| 75000 | 0.0075 | 0.84518 | 0.86558 | 1.2513 | 0.025853 | 0.068348 | 2.9134 | 0.19978 | 0.017274
| 80000 | 0.008 | 0.90152 | 0.84833 | 1.2264 | 0.026296 | 0.093065 | 2.9634 | 0.18389 | 0.0082499
| 85000 | 0.0085 | 0.95787 | 0.83029 | 1.2003 |  |  |  | SSE--> | 5.7379
| 90000 | 0.009 | 1.0142 | 0.81152 | 1.1732
| 95000 | 0.0095 | 1.0706 | 0.79207 | 1.1451
| 100000 | 0.01 | 1.1269 | 0.772 | 1.1161
| 105000 | 0.0105 | 1.1832 | 0.75137 | 1.0862
| 110000 | 0.011 | 1.2396 | 0.73023 | 1.0557
| 115000 | 0.0115 | 1.2959 | 0.70866 | 1.0245
| 120000 | 0.012 | 1.3523 | 0.68671 | 0.99276
| 125000 | 0.0125 | 1.4086 | 0.66444 | 0.96057
| 130000 | 0.013 | 1.465 | 0.64192 | 0.928
| 135000 | 0.0135 | 1.5213 | 0.6192 | 0.89516
| 140000 | 0.014 | 1.5777 | 0.59634 | 0.86212
| 145000 | 0.0145 | 1.634 | 0.57341 | 0.82897
| 150000 | 0.015 | 1.6904 | 0.55047 | 0.79579
| 155000 | 0.0155 | 1.7467 | 0.52756 | 0.76268
| 160000 | 0.016 | 1.803 | 0.50475 | 0.72971
| 165000 | 0.0165 | 1.8594 | 0.48209 | 0.69695
| 170000 | 0.017 | 1.9157 | 0.45963 | 0.66448
| 175000 | 0.0175 | 1.9721 | 0.43743 | 0.63238
| 180000 | 0.018 | 2.0284 | 0.41552 | 0.60071
| 185000 | 0.0185 | 2.0848 | 0.39397 | 0.56955
| 190000 | 0.019 | 2.1411 | 0.3728 | 0.53894
| 195000 | 0.0195 | 2.1975 | 0.35205 | 0.50895
| 200000 | 0.02 | 2.2538 | 0.33178 | 0.47964
| 205000 | 0.0205 | 2.3102 | 0.312 | 0.45105
| 210000 | 0.021 | 2.3665 | 0.29276 | 0.42324
| 215000 | 0.0215 | 2.4228 | 0.27408 | 0.39623
| 220000 | 0.022 | 2.4792 | 0.25599 | 0.37007
| 225000 | 0.0225 | 2.5355 | 0.2385 | 0.3448
| 230000 | 0.023 | 2.5919 | 0.22165 | 0.32043
| 235000 | 0.0235 | 2.6482 | 0.20544 | 0.297
| 240000 | 0.024 | 2.7046 | 0.1899 | 0.27453
| 245000 | 0.0245 | 2.7609 | 0.17502 | 0.25302
| 250000 | 0.025 | 2.8173 | 0.16082 | 0.2325

## ZimmPrep
Excess, Normalized Intensities (arbitrary units)
[5 empty] | c1/g.mL-1 | c2/g.mL-1 | c3/g.mL-1 | c4/g.mL-1 | c5/g.mL-1 | c6/g.mL-1 | c7/g.mL-1 | c8/g.mL-1
Detector | Type | Angle/deg | q/cm-1 | q2/cm-2 | 0.002601 | 0.004828 | 0.008176 | 0.010235 | 1e-05 | 1e-05 | 1e-05 | 1e-05 | <<Dummy values stored here toprevent computational overflows
1 | Wyatt | 22.5 | 53505.227 | 2862809314.2053 | 0.060165 | 0.1726 | 0.37578 | 0.504 | -0.084822 | -0.084822 | -0.084822 | -0.084822 | Replace with true concentrations, adjusted for loss to
2 | Wyatt | 28 | 66349.1956 | 4402215757.0604 | 0.059479 | 0.10331 | 0.20505 | 0.34592 | -0.076697 | -0.076697 | -0.076697 | -0.076697 | filter and/or centrifugation
3 | Wyatt | 32 | 75595.9554 | 5714748476.8025 | 0.088599 | 0.12572 | 0.20833 | 0.25143 | -0.035919 | -0.035919 | -0.035919 | -0.035919
4 | Blank | 38 | 89289.9136 | 7972688675.0404 | --- | --- | --- | --- | --- | --- | --- | ---
5 | Fiber/APD | 44 | 102739.1344 | 10555329732.2612 | 0.062634 | 0.17928 | 0.24498 | 0.24934 | 0.062634 | 0.062634 | 0.062634 | 0.062634
6 | Fiber/APD | 50 | 115906.7543 | 13434375692.5896 | 0.08358 | 0.093075 | 0.12064 | 0.11737 | 0.08358 | 0.08358 | 0.08358 | 0.08358
7 | Wyatt | 57 | 130864.963 | 17125638542.325 | 0.069606 | 0.097801 | 0.12776 | 0.15507 | -0.044936 | -0.044936 | -0.044936 | -0.044936
8 | Fiber/APD | 64 | 145334.9926 | 21122260085.1353 | 0.071988 | 0.078889 | 0.095937 | 0.10832 | 0.071988 | 0.071988 | 0.071988 | 0.071988
9 | Wyatt | 72 | 161205.2459 | 25987131315.9539 | 0.078214 | 0.10301 | 0.12495 | 0.1383 | -0.03243 | -0.03243 | -0.03243 | -0.03243
10 | Fiber/APD | 81 | 178116.807 | 31725596942.5504 | 0.065084 | 0.056925 | 0.063735 | 0.073856 | 0.065084 | 0.065084 | 0.065084 | 0.065084
11 | Wyatt | 90 | 193930.2188 | 37608929748.4912 | 0.075293 | 0.10074 | 0.12407 | 0.13574 | -0.031814 | -0.031814 | -0.031814 | -0.031814
12 | Fiber/APD | 99 | 208547.9862 | 43492262554.4321 | 0.095704 | 0.052966 | 0.055095 | 0.060668 | 0.095704 | 0.095704 | 0.095704 | 0.095704
13 | Blank | 108 | 221879.986 | 49230728181.0285 | --- | --- | --- | --- | --- | --- | --- | ---
14 | Fiber/APD | 117 | 233844.0219 | 54683026559.6783 | 0.09935 | 0.10311 | 0.11423 | 0.12293 | 0.09935 | 0.09935 | 0.09935 | 0.09935
15 | Wyatt | 126 | 244366.3316 | 59714904009.158 | 0.064786 | 0.091365 | 0.11379 | 0.13206 | -0.040699 | -0.040699 | -0.040699 | -0.040699
16 | Fiber/APD | 134 | 252456.5064 | 63734287602.2851 | 0.073718 | 0.15028 | 0.16713 | 0.18618 | 0.073718 | 0.073718 | 0.073718 | 0.073718
17 | Wyatt | 141 | 258527.6728 | 66836557617.2501 | 0.068348 | 0.094168 | 0.11999 | 0.14125 | -0.037971 | -0.037971 | -0.037971 | -0.037971
18 | Fiber/APD | 147 | 262964.6977 | 69150432220.4779 | 0.093065 | 0.085617 | 0.10062 | 0.11666 | 0.093065 | 0.093065 | 0.093065 | 0.093065
Excess Rayleigh factors in cm-1
[5 empty] | c1/g.mL-1 | c2/g.mL-1 | c3/g.mL-1 | c4/g.mL-1 | c5/g.mL-1 | c6/g.mL-1 | c7/g.mL-1 | c8/g.mL-1
Detector | Type | Angle/deg | q/cm-1 | q2/cm-2 | 0.002601 | 0.004828 | 0.008176 | 0.010235 | 1e-05 | 1e-05 | 1e-05 | 1e-05
1 | Wyatt | 22.5 | 53505.227 | 2862809314.2053 | 0.00019971 | 0.00057295 | 0.0012474 | 0.001673 | -0.00028156 | -0.00028156 | -0.00028156 | -0.00028156
2 | Wyatt | 28 | 66349.1956 | 4402215757.0604 | 0.00019744 | 0.00034292 | 0.00068064 | 0.0011483 | -0.00025459 | -0.00025459 | -0.00025459 | -0.00025459
3 | Wyatt | 32 | 75595.9554 | 5714748476.8025 | 0.0002941 | 0.00041731 | 0.00069153 | 0.00083461 | -0.00011923 | -0.00011923 | -0.00011923 | -0.00011923
4 | Blank | 38 | 89289.9136 | 7972688675.0404
5 | Fiber/APD | 44 | 102739.1344 | 10555329732.2612 | 0.00020791 | 0.0005951 | 0.00081321 | 0.00082768 | 0.00020791 | 0.00020791 | 0.00020791 | 0.00020791
6 | Fiber/APD | 50 | 115906.7543 | 13434375692.5896 | 0.00027744 | 0.00030896 | 0.00040046 | 0.00038961 | 0.00027744 | 0.00027744 | 0.00027744 | 0.00027744
7 | Wyatt | 57 | 130864.963 | 17125638542.325 | 0.00023105 | 0.00032465 | 0.00042409 | 0.00051475 | -0.00014916 | -0.00014916 | -0.00014916 | -0.00014916
8 | Fiber/APD | 64 | 145334.9926 | 21122260085.1353 | 0.00023896 | 0.00026187 | 0.00031846 | 0.00035958 | 0.00023896 | 0.00023896 | 0.00023896 | 0.00023896
9 | Wyatt | 72 | 161205.2459 | 25987131315.9539 | 0.00025963 | 0.00034195 | 0.00041477 | 0.0004591 | -0.00010765 | -0.00010765 | -0.00010765 | -0.00010765
10 | Fiber/APD | 81 | 178116.807 | 31725596942.5504 | 0.00021604 | 0.00018896 | 0.00021157 | 0.00024516 | 0.00021604 | 0.00021604 | 0.00021604 | 0.00021604
11 | Wyatt | 90 | 193930.2188 | 37608929748.4912 | 0.00024993 | 0.00033442 | 0.00041186 | 0.00045058 | -0.0001056 | -0.0001056 | -0.0001056 | -0.0001056
12 | Fiber/APD | 99 | 208547.9862 | 43492262554.4321 | 0.00031768 | 0.00017582 | 0.00018289 | 0.00020138 | 0.00031768 | 0.00031768 | 0.00031768 | 0.00031768
13 | Blank | 108 | 221879.986 | 49230728181.0285
14 | Fiber/APD | 117 | 233844.0219 | 54683026559.6783 | 0.00032979 | 0.00034228 | 0.0003792 | 0.00040805 | 0.00032979 | 0.00032979 | 0.00032979 | 0.00032979
15 | Wyatt | 126 | 244366.3316 | 59714904009.158 | 0.00021505 | 0.00030328 | 0.00037772 | 0.00043838 | -0.0001351 | -0.0001351 | -0.0001351 | -0.0001351
16 | Fiber/APD | 134 | 252456.5064 | 63734287602.2851 | 0.0002447 | 0.00049885 | 0.0005548 | 0.000618 | 0.0002447 | 0.0002447 | 0.0002447 | 0.0002447
17 | Wyatt | 141 | 258527.6728 | 66836557617.2501 | 0.00022688 | 0.00031259 | 0.0003983 | 0.00046888 | -0.00012604 | -0.00012604 | -0.00012604 | -0.00012604
18 | Fiber/APD | 147 | 262964.6977 | 69150432220.4779 | 0.00030892 | 0.0002842 | 0.000334 | 0.00038726 | 0.00030892 | 0.00030892 | 0.00030892 | 0.00030892
Kc/Excess Rayleigh factor in mol.g-1
[5 empty] | c1/g.mL-1 | c2/g.mL-1 | c3/g.mL-1 | c4/g.mL-1 | c5/g.mL-1 | c6/g.mL-1 | c7/g.mL-1 | c8/g.mL-1 | kscale: | 5000 | mL/g
Detector | Type | Angle/deg | q/cm-1 | q2/cm-2 | 0.002601 | 0.004828 | 0.008176 | 0.010235 | 1e-05 | 1e-05 | 1e-05 | 1e-05
1 | Wyatt | 22.5 | 53505.227 | 2862809314.2053 | 1.0487e-06 | 6.7854e-07 | 5.2779e-07 | 4.9262e-07 | -2.8599e-09 | -2.8599e-09 | -2.8599e-09 | -2.8599e-09
2 | Wyatt | 28 | 66349.1956 | 4402215757.0604 | 1.0608e-06 | 1.1337e-06 | 9.6727e-07 | 7.1775e-07 | -3.1629e-09 | -3.1629e-09 | -3.1629e-09 | -3.1629e-09
3 | Wyatt | 32 | 75595.9554 | 5714748476.8025 | 7.1215e-07 | 9.3162e-07 | 9.5203e-07 | 9.8748e-07 | -6.7536e-09 | -6.7536e-09 | -6.7536e-09 | -6.7536e-09
4 | Blank | 38 | 89289.9136 | 7972688675.0404
5 | Fiber/APD | 44 | 102739.1344 | 10555329732.2612 | 1.0074e-06 | 6.5328e-07 | 8.0958e-07 | 9.9575e-07 | 3.873e-09 | 3.873e-09 | 3.873e-09 | 3.873e-09
6 | Fiber/APD | 50 | 115906.7543 | 13434375692.5896 | 7.5491e-07 | 1.2583e-06 | 1.644e-06 | 2.1154e-06 | 2.9024e-09 | 2.9024e-09 | 2.9024e-09 | 2.9024e-09
7 | Wyatt | 57 | 130864.963 | 17125638542.325 | 9.0646e-07 | 1.1975e-06 | 1.5524e-06 | 1.6011e-06 | -5.3984e-09 | -5.3984e-09 | -5.3984e-09 | -5.3984e-09
8 | Fiber/APD | 64 | 145334.9926 | 21122260085.1353 | 8.7647e-07 | 1.4846e-06 | 2.0673e-06 | 2.292e-06 | 3.3697e-09 | 3.3697e-09 | 3.3697e-09 | 3.3697e-09
9 | Wyatt | 72 | 161205.2459 | 25987131315.9539 | 8.0671e-07 | 1.1369e-06 | 1.5873e-06 | 1.7952e-06 | -7.4802e-09 | -7.4802e-09 | -7.4802e-09 | -7.4802e-09
10 | Fiber/APD | 81 | 178116.807 | 31725596942.5504 | 9.6945e-07 | 2.0574e-06 | 3.1119e-06 | 3.3617e-06 | 3.7272e-09 | 3.7272e-09 | 3.7272e-09 | 3.7272e-09
11 | Wyatt | 90 | 193930.2188 | 37608929748.4912 | 8.38e-07 | 1.1625e-06 | 1.5985e-06 | 1.8291e-06 | -7.625e-09 | -7.625e-09 | -7.625e-09 | -7.625e-09
12 | Fiber/APD | 99 | 208547.9862 | 43492262554.4321 | 6.5928e-07 | 2.2112e-06 | 3.5999e-06 | 4.0925e-06 | 2.5347e-09 | 2.5347e-09 | 2.5347e-09 | 2.5347e-09
13 | Blank | 108 | 221879.986 | 49230728181.0285
14 | Fiber/APD | 117 | 233844.0219 | 54683026559.6783 | 6.3508e-07 | 1.1358e-06 | 1.7362e-06 | 2.0197e-06 | 2.4417e-09 | 2.4417e-09 | 2.4417e-09 | 2.4417e-09
15 | Wyatt | 126 | 244366.3316 | 59714904009.158 | 9.739e-07 | 1.2819e-06 | 1.743e-06 | 1.88e-06 | -5.9604e-09 | -5.9604e-09 | -5.9604e-09 | -5.9604e-09
16 | Fiber/APD | 134 | 252456.5064 | 63734287602.2851 | 8.559e-07 | 7.7933e-07 | 1.1867e-06 | 1.3336e-06 | 3.2907e-09 | 3.2907e-09 | 3.2907e-09 | 3.2907e-09
17 | Wyatt | 141 | 258527.6728 | 66836557617.2501 | 9.2315e-07 | 1.2437e-06 | 1.653e-06 | 1.7577e-06 | -6.3886e-09 | -6.3886e-09 | -6.3886e-09 | -6.3886e-09
18 | Fiber/APD | 147 | 262964.6977 | 69150432220.4779 | 6.7798e-07 | 1.3679e-06 | 1.9711e-06 | 2.1282e-06 | 2.6066e-09 | 2.6066e-09 | 2.6066e-09 | 2.6066e-09
Detector | Type | Angle/deg |  | q2/1010cm-2 | q2/1010cm-2 + kscale·c >>
1 | Wyatt | 22.5 |  | 0.28628 | 13.2913 | 24.4263 | 41.1663 | 51.4613 | 0.33628 | 0.33628 | 0.33628 | 0.33628
2 | Wyatt | 28 |  | 0.44022 | 13.4452 | 24.5802 | 41.3202 | 51.6152 | 0.49022 | 0.49022 | 0.49022 | 0.49022
3 | Wyatt | 32 |  | 0.57147 | 13.5765 | 24.7115 | 41.4515 | 51.7465 | 0.62147 | 0.62147 | 0.62147 | 0.62147
4 | Blank | 38
5 | Fiber/APD | 44 |  | 1.0555 | 14.0605 | 25.1955 | 41.9355 | 52.2305 | 1.1055 | 1.1055 | 1.1055 | 1.1055
6 | Fiber/APD | 50 |  | 1.3434 | 14.3484 | 25.4834 | 42.2234 | 52.5184 | 1.3934 | 1.3934 | 1.3934 | 1.3934
7 | Wyatt | 57 |  | 1.7126 | 14.7176 | 25.8526 | 42.5926 | 52.8876 | 1.7626 | 1.7626 | 1.7626 | 1.7626
8 | Fiber/APD | 64 |  | 2.1122 | 15.1172 | 26.2522 | 42.9922 | 53.2872 | 2.1622 | 2.1622 | 2.1622 | 2.1622
9 | Wyatt | 72 |  | 2.5987 | 15.6037 | 26.7387 | 43.4787 | 53.7737 | 2.6487 | 2.6487 | 2.6487 | 2.6487
10 | Fiber/APD | 81 |  | 3.1726 | 16.1776 | 27.3126 | 44.0526 | 54.3476 | 3.2226 | 3.2226 | 3.2226 | 3.2226
11 | Wyatt | 90 |  | 3.7609 | 16.7659 | 27.9009 | 44.6409 | 54.9359 | 3.8109 | 3.8109 | 3.8109 | 3.8109
12 | Fiber/APD | 99 |  | 4.3492 | 17.3542 | 28.4892 | 45.2292 | 55.5242 | 4.3992 | 4.3992 | 4.3992 | 4.3992
13 | Blank | 108
14 | Fiber/APD | 117 |  | 5.4683 | 18.4733 | 29.6083 | 46.3483 | 56.6433 | 5.5183 | 5.5183 | 5.5183 | 5.5183
15 | Wyatt | 126 |  | 5.9715 | 18.9765 | 30.1115 | 46.8515 | 57.1465 | 6.0215 | 6.0215 | 6.0215 | 6.0215
16 | Fiber/APD | 134 |  | 6.3734 | 19.3784 | 30.5134 | 47.2534 | 57.5484 | 6.4234 | 6.4234 | 6.4234 | 6.4234
17 | Wyatt | 141 |  | 6.6837 | 19.6887 | 30.8237 | 47.5637 | 57.8587 | 6.7337 | 6.7337 | 6.7337 | 6.7337
18 | Fiber/APD | 147 |  | 6.915 | 19.92 | 31.055 | 47.795 | 58.09 | 6.965 | 6.965 | 6.965 | 6.965
Scroll Down!
All this on the left is filled from other worksheets.
Scroll Down!
All this on the left is filled from other worksheets.
Hint: for starters, try kscale values that make k·cmax about the same as q2max/1010
Normally, that will be kscale = 1000 to 10000.

## ForNonZimm
OVERWRITE WITH YOUR DATA USING FORMAT SHOWN IN THE EXAMPLE AT RIGHT | [8 empty] | EXAMPLE FROM ROTATING ARM INSTRUMENT
The name of the file | [8 empty] | A:RUN4157.GAL | The name of the file
start time | [8 empty] | 11:23 | start time
start time | [8 empty] | 1996-03-18 | start date
start time | [8 empty] | 0.75988 | end time
start time | [8 empty] | 1996-03-18 | end date
first label | [8 empty] | DOW 1683 POLYSTYRENE IN TOLUENE. . . . . . . . . . . . . . . . . . . . . . . . . . . first label
second label | [8 empty] | WIESLAW WORST DAY. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . second label
operator | [8 empty] | WS. . . . . . . . | operator
wavelength in Angstroms | [8 empty] | 4880 | wavelength in Angstroms
Rayleigh standard material | [8 empty] | TOLUENE. . . . . . . | Rayleigh standard material
refractive index of Rayleigh standard | [8 empty] | 1.495 | refractive index of Rayleigh standard
refractive index of solvent for the polymer solutions | [8 empty] | 1.495 | refractive index of solvent for the polymer solutions
m = volume exponent (see documentation) | [8 empty] | 1 | m = volume exponent (see documentation)
lowest temperature during measurement | [8 empty] | 24.9 | lowest temperature during measurement
highest temperature during measurement | [8 empty] | 25.1 | highest temperature during measurement
0 | [8 empty] | 1018.118 | dark count
0.03 | [8 empty] | 24133.97 | Rayleigh standard intensity = Istd
uncertainty of Istd | [8 empty] | 120.0893 | uncertainty of Istd
90 | [8 empty] | 90 | angle where Rayleigh standard was measured
16 | [8 empty] | 9 | number of angles = Nang
number of concentrations = Nconc; includes solvent as c=0 | [8 empty] | 9 | number of concentrations = Nconc; includes solvent as c=0
0 | [8 empty] | 0 | c1
0.002601 | [8 empty] | 0.0001305 | c2
0.0048277 | [8 empty] | 0.000261 | c3
0.0081757 | [8 empty] | 0.000654 | c4
0.010235 | [8 empty] | 0.001305 | c5
#DIV/0! | [8 empty] | 0.001957 | c6
#DIV/0! | [8 empty] | 0.00261 | c7
#DIV/0! | [8 empty] | 0.003914 | c8
#DIV/0! | [8 empty] | 0.005219 | c9
22.5 | [8 empty] | 20 | first angle; next row is intensity (c=0 & up) and then sintensity
0 | 0.15722 | 0.45104 | 0.98197 | 1.317 | -0.22165 | -0.22165 | -0.22165 | -0.22165 | 81428.38 97591.7 122109.7 162909.5 197399.6 252660.4 301346.8 345541.3 373001.3
0 | 0.0039304 | 0.011276 | 0.024549 | 0.032926 | -0.0055413 | -0.0055413 | -0.0055413 | -0.0055413 | 189.3396 1144.371 75.31445 854.0734 235.379 396.2553 522.5136 2446.77 1101.427
28 | [8 empty] | 30
0 | 0.12669 | 0.22005 | 0.43676 | 0.73682 | -0.16337 | -0.16337 | -0.16337 | -0.16337 | 55574.58 66837.77 84346.55 112079.5 136769.3 177428.5 214188.9 232087 261054
0 | 0.0031673 | 0.0055012 | 0.010919 | 0.018421 | -0.0040842 | -0.0040842 | -0.0040842 | -0.0040842 | 197.434 392.5005 324.2591 173.9281 403.8692 568.4637 619.549 392.1764 669.6632
32 | [8 empty] | 40
0 | 0.16719 | 0.23723 | 0.39313 | 0.47447 | -0.067781 | -0.067781 | -0.067781 | -0.067781 | 43341.26 51996.1 65401.65 87735.49 106350.9 138269.6 166658.1 179940.2 203643.3
0 | 0.0041798 | 0.0059309 | 0.0098283 | 0.011862 | -0.0016945 | -0.0016945 | -0.0016945 | -0.0016945 | 239.8559 261.1551 328.7555 248.4714 564.8671 214.7371 180.5043 537.835 635.0795
38 | [8 empty] | 50
0 | [8 empty] | 36311.72 43244.43 55113.06 73145.96 89001.34 116999.8 137346.4 150733.3 170081.7
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 78.70634 234.4586 90.72663 191.5141 199.758 266.8599 273.3365 371.6528 367.2414
44 | [8 empty] | 60
0 | 0.090165 | 0.25808 | 0.35267 | 0.35894 | 0.090165 | 0.090165 | 0.090165 | 0.090165 | 32257.47 38655.18 48411.66 62924.06 78320.7 102318.9 119509.9 132199.1 148991
0 | 0.0022541 | 0.006452 | 0.0088167 | 0.0089735 | 0.0022541 | 0.0022541 | 0.0022541 | 0.0022541 | 106.8938 102.5643 131.3934 215.5929 239.1338 440.27 302.6808 93.53061 286.8959
50 | [8 empty] | 75
0 | 0.10911 | 0.1215 | 0.15749 | 0.15322 | 0.10911 | 0.10911 | 0.10911 | 0.10911 | 28955.09 33923.45 42198.49 55891.61 68934.62 90674.13 105791.5 116571.6 131802.6
0 | 0.0027277 | 0.0030375 | 0.0039372 | 0.0038304 | 0.0027277 | 0.0027277 | 0.0027277 | 0.0027277 | 150.9799 122.7043 159.0369 253.9555 388.7278 110.6202 293.2361 286.2568 516.4243
57 | [8 empty] | 90
0 | 0.082996 | 0.11661 | 0.15233 | 0.1849 | -0.05358 | -0.05358 | -0.05358 | -0.05358 | 27126.31 32055.37 39665.98 51456.23 62408.91 84488.31 98041.87 105803.9 122512.3
0 | 0.0020749 | 0.0029154 | 0.0038084 | 0.0046226 | -0.0013395 | -0.0013395 | -0.0013395 | -0.0013395 | 95.50737 194.1532 109.4389 230.5697 135.4062 152.2081 149.2928 281.5163 146.4149
64 | [8 empty] | 110
0 | 0.080094 | 0.087772 | 0.10674 | 0.12052 | 0.080094 | 0.080094 | 0.080094 | 0.080094 | 29638.82 34479.8 42685.7 54635.88 67755.86 89103.03 104593 114505.7 130324.7
0 | 0.0020024 | 0.0021943 | 0.0026685 | 0.003013 | 0.0020024 | 0.0020024 | 0.0020024 | 0.0020024 | 96.0753 157.568 165.4995 108.4058 114.0652 86.36407 222.5477 229.0522 273.5556
72 | [8 empty] | 120
0 | 0.082239 | 0.10831 | 0.13138 | 0.14542 | -0.034099 | -0.034099 | -0.034099 | -0.034099 | 32521.76 36138.69 46225.73 59356.04 73573.95 95478.31 110844.5 122913.1 140521
0 | 0.002056 | 0.0027079 | 0.0032845 | 0.0036356 | -0.00085247 | -0.00085247 | -0.00085247 | -0.00085247 | 170.4942 112.1322 216.0858 107.3936 102.3973 236.4535 258.8731 319.9878 372.4431
81
0 | 0.065895 | 0.057634 | 0.06453 | 0.074777 | 0.065895 | 0.065895 | 0.065895 | 0.065895
0 | 0.0016474 | 0.0014409 | 0.0016132 | 0.0018694 | 0.0016474 | 0.0016474 | 0.0016474 | 0.0016474
90
0 | 0.075293 | 0.10074 | 0.12407 | 0.13574 | -0.031814 | -0.031814 | -0.031814 | -0.031814
0 | 0.0018823 | 0.0025186 | 0.0031019 | 0.0033935 | -0.00079535 | -0.00079535 | -0.00079535 | -0.00079535
99
0 | 0.096897 | 0.053627 | 0.055782 | 0.061424 | 0.096897 | 0.096897 | 0.096897 | 0.096897
0 | 0.0024224 | 0.0013407 | 0.0013945 | 0.0015356 | 0.0024224 | 0.0024224 | 0.0024224 | 0.0024224
108
0
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
117
0 | 0.1115 | 0.11573 | 0.12821 | 0.13797 | 0.1115 | 0.1115 | 0.1115 | 0.1115
0 | 0.0027876 | 0.0028932 | 0.0032052 | 0.0034491 | 0.0027876 | 0.0027876 | 0.0027876 | 0.0027876
126
0 | 0.08008 | 0.11293 | 0.14065 | 0.16324 | -0.050307 | -0.050307 | -0.050307 | -0.050307
0 | 0.002002 | 0.0028233 | 0.0035163 | 0.004081 | -0.0012577 | -0.0012577 | -0.0012577 | -0.0012577
134
0 | 0.10248 | 0.20892 | 0.23234 | 0.25881 | 0.10248 | 0.10248 | 0.10248 | 0.10248
0 | 0.002562 | 0.0052229 | 0.0058086 | 0.0064704 | 0.002562 | 0.002562 | 0.002562 | 0.002562
141
0 | 0.10861 | 0.14963 | 0.19066 | 0.22445 | -0.060336 | -0.060336 | -0.060336 | -0.060336
0 | 0.0027151 | 0.0037409 | 0.0047666 | 0.0056113 | -0.0015084 | -0.0015084 | -0.0015084 | -0.0015084
147
0 | 0.17087 | 0.1572 | 0.18475 | 0.2142 | 0.17087 | 0.17087 | 0.17087 | 0.17087
0 | 0.0042718 | 0.00393 | 0.0046186 | 0.0053551 | 0.0042718 | 0.0042718 | 0.0042718 | 0.0042718
Written for a rotating arm instrument, the NonZimm.exe program subtracts solvent intensities and applies a sin(theta) correction for the scattering volume. We prepare SMALS data for use in NonZimm by undoing the sin(theta) correction and letting the solvent values all be zero. NonZimm is normally fed uncertainties, too. We take those as 2.5% of our normalized intensities. We take c1 = 0, as in NonZimm. Other concentrations are automatically entered as g/mL (see the ConcData worksheet) but you can override these (so they're colored pink instead of red).
<<< Usually, the intensity of toluene at 90o; probably parameters worksheet, G4
<<< Extremely unlikely, but you might change it. So, cell colored pink.