## Development notes
Developed May-June 2022 by Paul Russo

## Instructions
Purpose:
This workbook permits us to normalize detectors to a standard of finite size.
Using the normalization, you can then push ahead to measure an Rg using Guinier
Principle:
(there is no such thing as a perfect Rayleigh scatterer). Typically, we might use
a small latex particle, bovine serum albumen, or Brij-97 microemulsion.
The idea is something small that scatters a lot, well above solvent level and any
stray light.
The instrument has two kinds of detectors: Wyatt pin photodiodes plus
avalanche photodiode light sensors.
The Wyatt detectors can most easily be read from the front panel display.
The unit of measurement is probably in volts, but it does not matter.
The APD detectors measure in kHz, a totally different thing.
The objective is to figure out what factor puts the measured intensities on the
correct P(q) curve. We calculate the curve from the low-order P(q) expansion, 1-q2Rg2/3.
That will be good enough for any sensible normalization standard, but the equation
could be replaced with a detailed expression for, say, a larger latex sphere.
To use:
1 | Make a DLS run on your normalization standard, save as .ASC file
2 | Normally, this will give you a cumulants Gamma value and Rh value.
3 | Drop those data into the green-shaded cells at the top right; the results are copied where they need to go.
4 | Read the ALV's red panel display for all the laser and angles by gently pressing the selection button; type the results into the appropriate cells (color-coded).
button; type the results into the appropriate cells.
5 | Repeat the previous 2 steps for the solvent (likely a much weaker scatterer).
6 | Enter the Rh value from the DLS run.
7 | Enter your name and date.
8 | Enter the type of normalization standard.
9 | At this point, your normalization constants should be ready.
10 | You can then measure something known, preferably in a solvent with the same refractive index (subtract the solvent).
You can include the Rh value from DLS, but be careful about single-angle DLS (see Gather program for multiangle)
Apply the normalization factors to the intensities.
A Guinier plot or similar should give you the size expected for the known (if it's a sphere, don't forget to convert the DLS results to Rg).
11 | Then you can get Rg/Rh for shape analysis (e.g., should be sqrt(3/5) for solid, uniform spheres)
Testing
"Circular logic" tests with Brij microemulsion suggest the normalization factors were determined well enough to feed back into
the measured intensities and come up with smallish errors that might trace to the validity of Rh.
measurement in DLS (i.e., given the viscosity of the microemulsion, the true Rh of Brij blobs might be
smaller than measured.
Thus, Brij is probably perfect for getting the fudge factors right, but agreement between SLS and DLS might be imperfect.

## Normalize
Date of normalization |  | 2022-05-26
Normalizing standard |  | Brij microemulsion
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | y | 0.001 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 5.5 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 4.2603 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 93.6432
| LC |  |  | 79.2 |  | qmaxRg | 0.099666
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99669
Normalization
Comment¯ | Port # | Angle/degrees |  | ---- | Normalizing solution |  | Solvent background |  | Signal | P(q) | Factor
| Zero-ref | 0 |  | Detector type | Wyatt Reading | ALV Reading/kHz | Wyatt Reading | ALV Reading |  | 1 | 1 | Drop solution count rates |  | Drop solvent count rates
| 1 | 22.5 |  | Wyatt diode | 1.08 |  | 0 |  | 1.08 | 0.99986 | 0.9258 | here from .ASC file |  | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.72 |  | 0 |  | 0.72 | 0.99978 | 1.3886 | MeanCR1 [kHz] : | 763.5186 | MeanCR1 [kHz] : | 0
| 3 | 32 |  | Wyatt diode | 0.901 |  | 0 |  | 0.901 | 0.99972 | 1.1096 | MeanCR2 [kHz] : | 707.8195 | MeanCR2 [kHz] : | 0
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | --- | MeanCR3 [kHz] : | 680.9047 | MeanCR3 [kHz] : | 0
| 5 | 44 |  | Fiber/APD |  | 763.5186 |  | 0 | 763.5186 | 0.99947 | 0.001309 | MeanCR4 [kHz] : | 680.111 | MeanCR4 [kHz] : | 0
| 6 | 50 |  | Fiber/APD |  | 707.8195 |  | 0 | 707.8195 | 0.99931 | 0.0014118 | MeanCR5 [kHz] : | 163.3182 | MeanCR5 [kHz] : | 0
| 7 | 57 |  | Wyatt diode | 1.23 |  | 0 |  | 1.23 | 0.99911 | 0.81228 | MeanCR6 [kHz] : | 333.0665 | MeanCR6 [kHz] : | 0
| 8 | 64 |  | Fiber/APD |  | 680.9047 |  | 0 | 680.9047 | 0.99888 | 0.001467 | MeanCR7 [kHz] : | 417.5829 | MeanCR7 [kHz] : | 0
| 9 | 72 |  | Wyatt diode | 1.117 |  | 0 |  | 1.117 | 0.99858 | 0.89398 | MeanCR8 [kHz] : | 567.7117 | MeanCR8 [kHz] : | 0
| 10 | 81 |  | Fiber/APD |  | 680.111 |  | 0 | 680.111 | 0.9982 | 0.0014677
| 11 | 90 |  | Wyatt diode | 1.003 |  | 0 |  | 1.003 | 0.99778 | 0.99479
| 12 | 99 |  | Fiber/APD |  | 163.3182 |  | 0 | 163.3182 | 0.99731 | 0.0061066
| 13 | 108 |  | Empty | --- |  | --- |  | --- | --- | ---
| 14 | 117 |  | Fiber/APD |  | 333.0665 |  | 0 | 333.0665 | 0.99625 | 0.0029911
| 15 | 126 |  | Wyatt diode | 1.285 |  | 0 |  | 1.285 | 0.99565 | 0.77482
| 16 | 134 |  | Fiber/APD |  | 417.5829 |  | 0 | 417.5829 | 0.99508 | 0.0023829
| 17 | 141 |  | Wyatt diode | 0.709 |  | 0 |  | 0.709 | 0.99455 | 1.4027
| 18 | 147 |  | Fiber/APD |  | 567.7117 |  | 0 | 567.7117 | 0.99407 | 0.001751

## Measure Unknown
Date of normalization |  | 2022-05-26
Normalizing standard
Operator |  | Russo
Comment¯ | Monitors
| Front | beam stop (y/n) | y | 0.001 |  | Parameter
| p1 |  |  | 4.98 |  | Rh | 5.5 | nm
| p2 |  |  | 15.04 |  | Rg (sphere) | 4.2603 | nm
| p3 |  |  | 14.89 |  | lo/(n*DLS radius) | 93.6432
| LC |  |  | 79.2 |  | qmaxRg | 0.099666
| L |  |  | 4.92 |  | 1-(qmaxRg)2/3 | 0.99669
Normalized
Comment¯ | Port # | Angle/degrees |  | ---- | Normalizing solution |  | Solvent background |  | Signal | Intensity
| Zero-ref | 0 |  | Detector type | Wyatt Reading | ALV Reading/kHz | Wyatt Reading | ALV Reading |  | 1 | Drop solution count rates |  | Drop solvent count rates
| 1 | 22.5 |  | Wyatt diode | 1.08 |  | 0 |  | 1.08 | 0.99986 | here from .ASC file |  | here from .ASC file
| 2 | 28 |  | Wyatt diode | 0.72 |  | 0 |  | 0.72 | 0.99978 | MeanCR1 [kHz] : | 763.5186 | MeanCR1 [kHz] : | 0
| 3 | 32 |  | Wyatt diode | 0.901 |  | 0 |  | 0.901 | 0.99972 | MeanCR2 [kHz] : | 707.8195 | MeanCR2 [kHz] : | 0
| 4 | 38 |  | Empty | --- |  | --- |  | --- | --- | MeanCR3 [kHz] : | 680.9047 | MeanCR3 [kHz] : | 0
| 5 | 44 |  | Fiber/APD |  | 763.5186 |  | 0 | 763.5186 | 0.99947 | MeanCR4 [kHz] : | 680.111 | MeanCR4 [kHz] : | 0
| 6 | 50 |  | Fiber/APD |  | 707.8195 |  | 0 | 707.8195 | 0.99931 | MeanCR5 [kHz] : | 163.3182 | MeanCR5 [kHz] : | 0
| 7 | 57 |  | Wyatt diode | 1.23 |  | 0 |  | 1.23 | 0.99911 | MeanCR6 [kHz] : | 333.0665 | MeanCR6 [kHz] : | 0
| 8 | 64 |  | Fiber/APD |  | 680.9047 |  | 0 | 680.9047 | 0.99888 | MeanCR7 [kHz] : | 417.5829 | MeanCR7 [kHz] : | 0
| 9 | 72 |  | Wyatt diode | 1.117 |  | 0 |  | 1.117 | 0.99858 | MeanCR8 [kHz] : | 567.7117 | MeanCR8 [kHz] : | 0
| 10 | 81 |  | Fiber/APD |  | 680.111 |  | 0 | 680.111 | 0.9982
| 11 | 90 |  | Wyatt diode | 1.003 |  | 0 |  | 1.003 | 0.99778
| 12 | 99 |  | Fiber/APD |  | 163.3182 |  | 0 | 163.3182 | 0.99731
| 13 | 108 |  | Empty | --- |  | --- |  | --- | ---
| 14 | 117 |  | Fiber/APD |  | 333.0665 |  | 0 | 333.0665 | 0.99625
| 15 | 126 |  | Wyatt diode | 1.285 |  | 0 |  | 1.285 | 0.99565
| 16 | 134 |  | Fiber/APD |  | 417.5829 |  | 0 | 417.5829 | 0.99508
| 17 | 141 |  | Wyatt diode | 0.709 |  | 0 |  | 0.709 | 0.99455
| 18 | 147 |  | Fiber/APD |  | 567.7117 |  | 0 | 567.7117 | 0.99407

## Guinier
Normalized
Port | Angle/degrees | q/cm-1 | q2/1010cm-2 | q2/10-4nm-2 | intensity, I | ln (I)
1 | 22.5 | 47599.9678 | 0.22658 | 0.22658 | 0.99986 | -6.0308e-05 | 0.99986 | -6.0308e-05
2 | 28 | 59026.3747 | 0.34841 | 0.34841 | 0.99978 | -9.3398e-05 | 0.99978 | -9.3398e-05
3 | 32 | 67252.5891 | 0.45229 | 0.45229 | 0.99972 | -0.00012199 | 0.99972 | -0.00012199
5 | 44 | 91400.0326 | 0.8354 | 0.8354 | 0.99947 | -0.00023067 | 0.99947 | -0.00023067
6 | 50 | 103114.37 | 1.0633 | 1.0633 | 0.99931 | -0.0002979 | 0.99931 | -0.0002979
7 | 57 | 116421.6728 | 1.3554 | 1.3554 | 0.99911 | -0.00038719 | 0.99911 | -0.00038719
8 | 64 | 129294.6757 | 1.6717 | 1.6717 | 0.99888 | -0.00048818 | 0.99888 | -0.00048818
9 | 72 | 143413.3626 | 2.0567 | 2.0567 | 0.99858 | -0.00061795 | 0.99858 | -0.00061795
10 | 81 | 158458.4303 | 2.5109 | 2.5109 | 0.9982 | -0.00078223 | 0.9982 | -0.00078223
11 | 90 | 172526.5491 | 2.9765 | 2.9765 | 0.99778 | -0.00096593 | 0.99778 | -0.00096593
12 | 99 | 185530.9844 | 3.4422 | 3.4422 | 0.99731 | -0.001169 | 0.99731 | -0.001169
14 | 117 | 208035.1499 | 4.3279 | 4.3279 | 0.99625 | -0.0016337 | 0.99625 | -0.0016337
15 | 126 | 217396.1344 | 4.7261 | 4.7261 | 0.99565 | -0.0018952 | 0.99565 | -0.0018952
16 | 134 | 224593.4137 | 5.0442 | 5.0442 | 0.99508 | -0.0021442 | 0.99508 | -0.0021442
17 | 141 | 229994.5183 | 5.2897 | 5.2897 | 0.99455 | -0.0023747 | 0.99455 | -0.0023747
18 | 147 | 233941.8381 | 5.4729 | 5.4729 | 0.99407 | -0.0025817 | 0.99407 | -0.0025817
Intercept
[8 empty] | Slope | 0.00028
[8 empty] | Rg,Guinier | 2.8983 | nm
[8 empty] | R if sphere | 3.7417 | nm
[8 empty] | Rh from DLS | 5.5 | nm

## Parameters
Solvent | Water
Wavelength/nm | 685
Wavelength/cm | 6.85e-05
Refractive index | 1.33